Information Sheet

USING AUTOCLAVES

Pressure steam sterilizers (autoclaves) are used in laboratories both for sterilisation of media and equipment required for the culture of microorganisms, and for sterilisation of discarded cultures and waste materials. Pressure steam sterilisers operate at high pressures and temperatures, and appropriate measures must be taken for personnel safety.

Autoclaves utilise moist heat, in the form of saturated steam under pressure, to destroy microbial life. Steam sterilisation is the most reliable sterilisation method.

The autoclave is used to sterilise:

- reusable equipment after it has been disinfected and washed

- biological waste before disposal.

1.1	Autoclave effectiveness

Autoclaves must be used properly to effect successful steam sterilisation. The following elements all contribute to autoclave effectiveness.

- Temperature: The temperature inside the autoclave must be at least 121°C.

- Time: The time required for sterilisation varies, but it will never be less than 30 minutes. The time is measured from the point at which the temperature of the material to be sterilised reaches the required autoclaving temperature - the tighter the autoclave is packed, the longer it will take for the material in the centre of the load to reach the temperature required.

- Contact: Saturated steam must contact all areas of the load. Sterilisation will fail if:

- air pockets or inadequate steam supply prevent steam saturation

- the load exceeds the capacity of the autoclave so that steam cannot contact all areas of the load.

1.2	General guidelines

There are a number of general procedures that should be followed when using autoclaves to ensure effective sterilisation.

- Indicators: Indicators, such as autoclave tape, indicate with a visible colour change that they have been through the autoclave process. At the end of the autoclave cycle, indicators should be checked to ensure they have changed colour.

- Biological indicators: There are commercially available test kits that use bacterial spores to test the autoclave efficiency. Autoclaves should be tested using biological indicators at regular intervals.

Information Sheet

USING AUTOCLAVES (Continued)

- Overfilling

  - Steam and heat cannot easily penetrate a densely packed autoclave bag. If an autoclave is overfilled the outer contents of the bag will be sterilised, but the innermost part will be unaffected.

  - An overpacked autoclave chamber does not allow efficient steam distribution, and so sterilisation efficiencies will be reduced.

- Monitoring: Autoclave temperature, pressure and cycle duration time should be monitored during each cycle. (Some autoclaves have charts that trend the temperatures and pressures inside the autoclave chamber throughout each cycle).

- Maintenance: Autoclaves must be covered by a regular preventative maintenance program that is performed by a technician certified by the manufacturer.

1.3	Autoclave bags

Most items to be autoclaved are packed into autoclave bags, which come in a variety of types and sizes. Paper autoclave bags are commonly used when autoclaving reusable equipment after it has been disinfected and washed. These bags are permeable to steam.

The following guidelines should be followed when using paper autoclave bags.

- Reusable items that have been autoclaved should be packed away in their autoclave bags so that the equipment inside remains sterile.

- Paper autoclave bags should be closed by folding down and taping the open end. Plastic biohazard bags are commonly used when autoclaving biological waste before disposal.

1.3.1	The following guidelines should be followed when autoclaving with plastic bags.

- Autoclave bags are commonly put into plastic or stainless steel pans to catch any liquids that may drain out of the autoclave bag.

- Autoclave bags should be left partially open, to allow steam to penetrate the bag.

- Where possible, water should be added to autoclave bags to facilitate saturated steam contact. However, water should not be added if by doing so there is a chance that biohazardous materials may splash out of the bag.

1.3.2	Autoclaving without autoclave bags

- Some reusable equipment is autoclaved without autoclave bags, and in this instance the following guidelines should be followed.

Information Sheet

USING AUTOCLAVES (Continued)

- Tubular equipment such as pipettes, and equipment such as conical flasks should be plugged with a wad of cotton wool, so that the interior of the equipment remains sterile after autoclaving.

- Small bottles and tubes should be loaded in open mesh baskets.

- Lids or caps should be loosened.

1.4	Autoclave safety

Like other laboratory equipment, autoclaves must be used in accordance with appropriate safety procedures. Safety precautions that should be followed when using autoclaves include the following.

- Wear personal protective equipment, including heat-resistant gloves, safety glasses and a laboratory coat when operating an autoclave.

- Be careful when opening the autoclave door. Opening the autoclave door too quickly can result in glassware breakage and steam burns.

- Allow steam to escape from the autoclave before attempting to remove the contents of the autoclave.

- Never autoclave sealed containers - this could cause an explosion inside the autoclave. Large bottles with narrow necks can simulate sealed containers if filled with too much liquid.

- Never autoclave solvents, volatile or corrosive chemicals, or any radioactive materials.

- If there is a spill inside an autoclave, allow the autoclave to cool before attempting to clean up the spill.

- If glass breaks in the autoclave use tongs, forceps or other suitable tools to clear it away. Do not use bare or gloved hands to pick up broken glassware.

- Do not put sharp or pointed contaminated objects into an autoclave bag, as this could cause a 'needle stick injury'. Instead, place sharp or pointed objects in an appropriate rigid sharps disposal container.

- Be careful when handling an autoclave bag full of infectious waste, in case sharp objects have been inadvertently placed in the bag. Never lift an autoclave bag from the bottom, instead handle it from the top.

Using an autoclave

Pressure: 	Sterilising is done according to steam pressure and this is determined by the temperature.

Temperature:	The pressure setting (i.e. temperature) for a run is dependent upon the temperature sensitivity of the items being autoclaved.

Time: 	The length of the run depends on the pressure chosen and the amount of load and /or the volume in individual bottles.

Exhaust: 	The jacket can be set to slow or fast exhaust or fast and dry.

Slow for agars, liquids, medias etc;

Fast for dry materials, rubbish/waste, unsealed bottles etc.  Note: a fast exhaust will cause liquids to bubble vigorously and may cause liquid loss.

Fast and dry for clean pipettes, glassware, forceps, packets etc. but NOT for plastics.

The time and pressure depend on the contents of the run so it is important not to mix items/media with greatly different temperature ranges in the same run.

Information Sheet

USING AUTOCLAVES (Continued)

Common pressures and times of autoclave runs

| Temp. | Pressure | Time | Exhaust | Applications |
| --- | --- | --- | --- | --- |
| 121OC | 15psi | 15 min | Slow | Liquids / agar |
| 121OC | 15psi | 20 min | Slow | Large volumes of liquid/agar |
| 121OC | 15psi | 20 min | Fast | Rubbish (no liquids) |
| 109OC | 5psi | 45 min | Slow | Heat sensitive media |
| 115OC | 10psi | 15 min | Slow | Heat sensitive media |
| 118OC | 12psi | 10 min | Slow | Heat sensitive media |
| 134OC | 30psi | 4 min | Fast | Glass / equipment |

Decontamination method for Prion Contaminated Material

Prions are the cause of scrapie and other transmissible spongiform encephalopathies, they produce slow, progressive and fatal diseases of the central nervous system.

Prions are resistant to most traditional methods of inactivation used for other microorganisms. Because of the difficulties in inactivating the infectivity, prions pose particular problems in the laboratory. Prions should only be handled in dedicated laboratories using dedicated equipment.

Current recommendations for the sterilisation of articles or specimens that could be contaminated by prions are 18 minutes at 134°C in a pre-vacuum pressure steam steriliser or 1 hour at 132°C in a downward displacement pressure steam steriliser. The recommended chemical disinfectant for effective decontamination of articles or specimens that could be contaminated by prions is 2% sodium hypochlorite for 1 hour. 

Current recommendations for decontaminating heat-resistant instruments that may be contaminated by prions involves a combination of pressure steam sterilising and chemical treatment. This can be either:
a) immerse in 1M sodium hydroxide and heat in a gravity displacement autoclave at 121°C for 30 minutes, clean, rinse in water then subject to routine sterilisation (refer to pressure steam sterilising section); or
b) immerse in 1M sodium hydroxide or 2% sodium hypochlorite for 1 hour, rinse, transfer instruments to water. Heat in a gravity displacement autoclave at 121°C for 1 hour, clean and subject to routine sterilisation